## System_Suitability
Parameter | Observed value | Limits | Review Decision
No. of theoretical plates | 1875 | > 1000
Tailing factor (symmetry) | 1.37 | < 1.5
Regression equation | y = 5.22229e-007x + 0.460672
Correlation coefficient (r2) | 0.99994 | ∼1 (0.99–1.00)

## Accuracy
Metformin Concentration (µg/mL) | Mean (µg/mL) | Recovery (%) | SD | RSD (%)
2.5 | 2.466 | 98.656 | 0.061 | 2.485
5 | 4.943 | 98.857 | 0.044 | 0.887
10 | 10.036 | 100.362 | 0.031 | 0.305
15 | 15.077 | 100.513 | 0.086 | 0.57
20 | 20.201 | 101.003 | 0.107 | 0.531
40 | 40.035 | 100.086 | 0.236 | 0.589

## Robustness
Run | Parameter | Condition | Peak_Area_Mean | Peak_Area_SD | Peak_Area_RSD_Percent | RT_Mean | RT_SD | RT_RSD_Percent
1 | Column temperature (°C) | 30 | 56824654 | 268221.6 | 0.472 | 1.192 | 0.0011 | 0.0919
2 | Flow rate (mL/min) | 0.9 | 50742178 | 213868.2 | 0.4215 | 1.0677 | 0.0008 | 0.0765
3 | Injection volume (µL) | 18 | 51373063 | 233477.1 | 0.4545 | 1.202 | 0.0006 | 0.0526
4 | Mobile phase pH | 3.5 | 56675050 | 257438.9 | 0.4542 | 1.2 | 0.0006 | 0.0527

## QC_Review
Parameter | Result | Acceptance Criteria | QC Decision | Reviewer Comments | QC Summary
Theoretical Plates | 1875 | >1000 | PASS | Meets system suitability criteria | PASS | 15
Tailing Factor | 1.37 | <1.5 | PASS | Meets system suitability criteria | INVESTIGATE | 4
Linearity (R²) | 0.99994 | >=0.999895 | PASS | Meets the reported UHPLC linearity criterion. | TOTAL REVIEWED | 19
Accuracy – Recovery (2.5) | 98.656 | 90–110% | PASS | Recovery 98.656% is within the acceptance range. | PASS Rate | 78.95%
Precision – %RSD (2.5) | 2.485 | <10% | PASS | RSD 2.485% is below the reported 10% limit.
Accuracy – Recovery (5) | 98.857 | 90–110% | PASS | Recovery near 100%; acceptance criteria need verification
Precision – %RSD (5) | 0.887 | <10% | PASS | - RSD 0.887% is below the reported 10% limit.
Accuracy – Recovery (10) | 100.362 | 90–110% | PASS | - Recovery 100.362% is within the reported acceptance range.
Accuracy – Recovery (15) | 100.513 | 90–110% | PASS | Recovery 100.513% is within the reported acceptance range.
Accuracy – Recovery (20) | 101.003 | 90–110% | PASS | Recovery 101.003% is within the reported acceptance range.
Accuracy – Recovery (40) | 100.086 | 90–110% | PASS | - Recovery 100.086% is within the reported acceptance range.
Precision – %RSD (10) | 0.305 | <10% | PASS | - RSD 0.305% is below the reported 10% limit.
Precision – %RSD (15) | 0.57 | <10% | PASS | RSD 0.570% is below the reported 10% limit
Precision – %RSD (20) | 0.531 | <10% | PASS | RSD 0.531% is below the reported 10% limit.
Precision – %RSD (40) | 0.589 | <10% | PASS | RSD 0.589% is below the reported 10% limit.
Robustness – Column Temperature (30°C) | 56824654 | No numeric limit reported | INVESTIGATE | Peak area RSD 0.472%; no numeric acceptance limit reported.
Robustness – Flow Rate (0.9 mL/min) | 50742178 | No numeric limit reported | INVESTIGATE | Mean peak area 50,742,178; no numeric acceptance limit reported.
Robustness – Injection Volume (18 µL) | 51373063 | No numeric limit reported | INVESTIGATE | Mean peak area 51,373,063; no numeric acceptance limit reported.
Robustness – Mobile Phase pH (3.5) | 56675050 | No numeric limit reported | INVESTIGATE | Mean peak area 56,675,050; no numeric acceptance limit reported.

## Final_Report
Metformin UHPLC – Analytical Data Review Report
Project Objective
Review published Metformin UHPLC validation data, evaluate reported results against documented criteria, and identify findings requiring further investigation.
QC Review Summary
Total Parameters Reviewed: 19
PASS: 15 (78.95%)
INVESTIGATE: 4 (21.05%)
Key Findings
15 results were classified as PASS based on the criteria used in this review.
Four robustness results require further investigation because numeric acceptance limits were not reported.
Reviewer Recommendations
Verify documented acceptance criteria for the four robustness results before assigning final PASS or FAIL decisions.
Data Source
Published Metformin UHPLC study (PMC11233872)
https://pmc.ncbi.nlm.nih.gov/articles/PMC11233872/